10 MHz Receiver

Kosh Yokota, Dick Plambeck  12/23/03

Revision A, 1/9/04

Revision B, 2/13/04

Revision C, 2/25/04

Function: generates 10 MHz reference signal and 1 pps clock tick at the antenna, from 10 MHz pulse train sent from lab via optical fiber.

Brief circuit description:  The laser signal is received by a Fermionics FD80FC photodiode.  The 10 MHz pulse train from the photodiode is amplified and used to phaselock a 10 MHz crystal oscillator (Frequency Management VC46).  The TTL output of the crystal oscillator is amplified, passed through a 10 MHz filter to produce a clean sine wave, then split to provide four 10 MHz outputs.   A sample of the incoming 10 MHz pulse train also is sent to a comparator and retriggerable one shot.  If a pulse is missing, the one shot generates a 1.5 microsecond wide clock tick, which is available as a positive-going TTL pulse and as an RJ485 differential signal.

Input:

- laser signal (1310 or 1550 nm) on single mode fiber, FC-APC connector.  The expected optical intensity is –16 dBm, or 25 microwatts.  The acceptable input level is –21 dBm or greater; the circuit functions down to about –27 dBm.

Output:

- Four SMA female connectors with 10 MHz sine wave, +8 dBm; note that crystal oscillator generates 10 MHz output even if it is not phaselocked

- One SMA female connector with 1 pps tick – 1.5 microseconds wide, positive going TTL.

- One RJ45 jack with RJ485 differential signals:

| 1 | 10 MHz TTL L |
| --- | --- |
| 2 | 10 MHz TTL H |
| 3 | 1 pps L |
| 4 | NC |
| 5 | NC |
| 6 | 1 pps H |
| 7 | NC |
| 8 | NC |

Indicator LEDs:

- optical pwr: 	green for > -21 dBm (acceptable operation)

yellow for  -27 dBm < pwr < -21 dBm (marginal operation)

off for pwr < -27 dBm

- 10 MHz lock:	green if phaselocked

Power supply/monitor through 9-pin D subminiature connector, male:

| 1 | -15 V, 50 mA |
| --- | --- |
| 2 | DS2401 serial number |
| 3 | pwr GND |
| 4 | gnd for photocurrent monitor |
| 5 | +15 V, 175 mA |
| 6 | gnd for DS2401 |
| 7 | photocurrent OK: 0=no, 1=yes (superfluous indicator) |
| 8 | photocurrent monitor, 10V/mA; expected level is 0.2 V |
| 9 | 10 MHz lock alarm, TTL: low if locked, high if unlocked |

1 pps tick as generated by missing pulse

Figure 1.  Top trace shows 1 pps clock tick output from SMA connector on 10 MHz receiver box.  Bottom trace shows the 10 MHz pulse train received from the lab transmitter over the fiberoptic link; a missing pulse encodes the 1 pps tick.

Thermal behavior

Figure 2.  Thermal test of 10 MHz receiver.  Top curve is 10 MHz phase, bottom curve is temperature.  The thermal coefficient is approximately –0.1 degree/C.